## Slide 1
IMC Installation steps
9 oktober 2013
Marco & Thomas

## Slide 2: Contents:
Pre-installation step 1 Positioning Gear Box
Pre-installation step 2 Positioning Mirror Vacuum clamp
Installation step 1 Positioning Mirror
Installation step 2 Wiring Mirror
Installation step 3 Positioning Reaction Mass
Installation step 4 Wiring Reaction Mass (RM)
Installation step 5 Transport to Tower
Installation step 6 Connection to Marionetta in Tower
Installation step 7 Mount Safety Stops and Baffles
Installation step 8 Remove Installation Frame from IMC
Installation complete

## Slide 3
Nomenclatures
RM Support
Mirror Support
Translation rail
Vacuum Pump
Battery for
Vacuum pump
Vacuum System
Mirror
Vacuum Clamp
Mounting plate
Gear box
Switch Vacuum pump

## Slide 4
Pre-installation step 1
Gear box
Actions:
Mount Gear box on top
Place RM support on rail
Mount RM on RM support
Check with 2 Plumb Lines paralism Gear Box with RM
Fix Gear Box with 4 bolts
Positioning Gear Box
Plump Line

## Slide 5
Pre-installation step 2
Actions:
Check position Mirror Vacuum Clamp with the Dummy Mirror
Fix position Mirror Vacuum Clamp
Positioning Mirror Vacuum clamp
Plump Line
Dummy Mirror
Mirror
Vacuum clamp

## Slide 6
Installation step 1
Actions:
a) Place Mirror Support on rail
Positioning Mirror

## Slide 7
Installation step 1b
Actions:
Place Mirror on Mirror Support
Positioning Mirror

## Slide 8
Installation step 1c
Actions:
Place the Top clamp just above the Mirror
Check horizontal position Mirror with wire spaces/magnets and the two Horizontal strips (pink)
Top clamp
Horizontal strip
Positioning Mirror

## Slide 9
Installation step 1d
Actions:
Place Klarez O-ring in Vacuum Clamp
Slide gently the Mirror with support close to the Vacuum clamp
Adjust the Vacuum clamp in the center and parallel to the surface of the Mirror
Start Vacuum pump
Slide gently the Mirror to the Vacuum clamp until the vacuum is closed (mirror touch the Kalrez O-ring)
Vacuum clamp
Positioning Mirror
Switch Vacuum Pump
Vacuum Pump
Klarez O-ring

## Slide 10
Installation step 1e
Positioning Mirror
Actions:
Move up the Vacuum clamp (with Mirror) by a known distance to get the Mirror Support free from the Mirror
Slide gently the Mirror clamp to the end of the Mirror.
Move down the Vacuum clamp.
Fix the Mirror Clamp on the rail
Now there is access for wiring and the Mirror is safe when the vacuum lost.

## Slide 11
Installation step 2
Actions:
Wire the Mirror with two CuBe Wires
Tension the two Wires in the Gear box. This tension will be adjust with a frequency probe used for the Multi-Payload Suspensions (not yet tested)
Wiring Mirror
Two CuBe Wires

## Slide 12
Installation step 3
Positioning Reaction Mass
RM
Dowel pin
Actions:
Place the RM support on the Translation rail
Place the RM in the RM Support, a dowel pin positioned and 4 bolts do the fixation
RM Support
Fixation Bolts

## Slide 13
Installation step 3b
Actions:
Slide gently the Recoil with Support to the Mirror
Position is correct when the wirespacres are in the center of the holes through the RM (both sides)
Positioning Reaction Mass

## Slide 14
Installation step 4
Wiring Reaction Mass (RM)
Actions:
Wire the RM with two CuBe/C85 Wires
Tension the two Wires in the Gear box. This tension will be adjusted with a frequency probe used for the Multi-Payload Suspensions

## Slide 15
Actions:
Secure the whole installation frame
Wrap the setup in plastic
.
Installation step 5
Transport to Tower

## Slide 16
Installation step 6
Remark:
Marionetta is already installed
Connection to Marionetta in Tower
Marionetta

## Slide 17
Installation step 6b
Actions:
IMC going up on Lift
Connection to Marionetta in Tower

## Slide 18
Installation step 6c
Actions:
Connect IMC Gear Box to Marionetta with 4 bolts
Connection to Marionetta in Tower
Bolt Connections
Note:
mirror and RM are suspended from wires, and are supported by supports

## Slide 19
Installation step 7
Mount Safety Stops and Baffles
Actions:
Remove Mirror Vacuum Clamp from the frame, to get access to front of RM.
remove Mirror Vacuum Clamp
Note:
mirror is freely hanging,
RM is suspended from wires, and is supported by supports

## Slide 20
Installation step 7b
Mount Safety Stops and Baffles
Actions:
Mount the 4 Safety Stops.
Safety Stop
Note:
mirror is freely hanging,
RM is suspended from wires, and is supported by supports

## Slide 21
Installation step 7c
Mount Safety Stops and Baffles
Actions:
Mount the Baffle Support Ring.
Baffle Support Ring

## Slide 22
Installation step 7d
Mount Safety Stops and Baffles
Actions:
Mount a Baffle.
Baffle

## Slide 23
Installation step 7e
Mount Safety Stops and Baffles
Actions:
Mount the last Baffle.

## Slide 24
Installation step 7e
Mount Safety Stops and Baffles
Note:
mirror is freely hanging,
RM is suspended from wires, and is supported by supports

## Slide 25
Installation step 8
Remove Installation Frame from IMC
Actions:
Disconnect the IMC from the Installation frame, by removing the two Gear box connection plates
Gear Box connection Plate

## Slide 26
Installation step 8
Remove Installation Frame from IMC
Actions:
remove the four bolts between
RM support and RM.
Note:
mirror is freely hanging,
RM is suspended from wires only and not connected to support

## Slide 27
Installation step 8b
Remove Installation Frame from IMC
Actions:
Lower Lift

## Slide 28
Installation step 8b
Remove Installation Frame from IMC

## Slide 29
Installation step 8b
Remove Installation Frame from IMC

## Slide 30
Installation complete