## Slide 1: SAT Electronics ADC-DAC Performances
Cascina , October 28th 2014
Alberto Gennai (INFN Pisa)
M. Bitossi
V. Boschi, C. Carissimi, A. Gennai, C. Magazzu, D. Passuello

## Slide 2: Board Testing

## Slide 3: Board Block Diagram

## Slide 4: DAC Noise and Distortion Measurement Setup
DAC
Analog notch filter
FFT Analyzer
Analog Front End
DSP

## Slide 5: DAC Total Harmonic Distortion
74.75 Hz

## Slide 6: DAC Noise 1/2
ADC Conversion Rate: 320 kSPS
Signal set at 0.9 x Full Scale (FS = 20Vpp)
100nV/sqrt(Hz)

## Slide 7: DAC Noise 2/2 (Low Frequency)
ADC Conversion Rate: 320 kSPS
Signal set at 0.9 x Full Scale (FS = 20Vpp)
Virgo/Virgo+ DAC Noise
1uV/sqrt(Hz)

## Slide 8: ADC Noise and Distortion Measurement Setup
DAC#1
FFT Analyzer
Analog Front End
DSP
Analog Front End
Analog Front End
DAC#2
ADC
1. DAC #1 generates a sine wave
2. DSP reads ADC, apply a digital notch filter and amplifies by a factor 1000 before sending output to DAC #2

## Slide 9: ADC-DSP-DAC Frequency Response
Group delay ~ 42 usec (~650 usec in Virgo+)
24 kHz

## Slide 10: ADC Noise 1/2 (High Frequency)
ADC Conversion Rate: 3.84 MSPS
12 x decimation  320 kSPS
0 Vdc output (FS = 10Vpp)
100nV/sqrt(Hz)
70nV/sqrt(Hz)

## Slide 11: ADC Noise 2/2 (Low Frequency)
ADC Conversion Rate: 3.84 MSPS
12 x decimation  320 kSPS
Signal set at 0.8 x Full Scale (FS = 10Vpp)
2uV/sqrt(Hz)

## Slide 12: ADC THD
ADC Conversion Rate: 3.84 MSPS
12 x decimation  320 kSPS
Signal set at 0.8 x Full Scale (FS = 10Vpp)
THD = -96.8 dB
74.75 Hz