## Slide 1: ENV training Part III Data analysis techniques
Irene Fiori on behalf of the Commissioning team
VIR-0948A-19
Virgo training sessions
September 26 2019

## Slide 2: Outline of Part III
Inspection of Time-Frequency plots
Correlations
  On/off lines
  Moving lines
  Brute force approach
Coherence (BruCo)
Line DB and NoeMi Line finder
Scattered light hunting
  Spectral features (How to recognize you are dealing with S.L)
  Investigation techniques

## Slide 3: Inspection of Time-Frequency images
The VIM (see Nicolas Arnaud’s talk)

## Slide 4: Inspection of Time-Frequency images
Look for:
  Lines or bumps that
  Start/stop/move
  Periodic patterns
  Combs (equally spaced lines)
  Sidebands

## Slide 6: Inspection of Time-Frequency images
Browse through VIM spectrogram and Look for correlated patterns
Where:
  Hrec and DARM
  LSC_PRCL/MICH  orthogonal dofs
  SNEB_B7/SWEB_B8/SPRB_B4 power of peak-off beams
  SSFS_Err  frequency noise
  BsX_*  input beam jitter
  ENV_*SEIS, *MAG, *MIC, *UPS, *IPS  source is in the environment
  …. all VIM channels are interesting, have been selected on purpose.
Did you know?
You can reproduce all VIM spectra and zoom/play as you like
https://logbook.virgo-gw.eu/virgo/?r=38996
USEFUL to inspect LONG time stretches
elog 37121
BsX
CEB_IPS

## Slide 7: Correlate on/off lines
On/off lines usually means some device turned on/off …
Asked VACUUM guys if they did something…. “well… maybe we switched off one PLC driver”. YES!
Atomic GPS receiver board was swapped
Often NEW NOISE appear during MAINTENANCES
Always suspect of “transparent activities” 
RF noise at 1200Hz around 6MHz

## Slide 8: Elog 41624 “Forest of lines moving yesterday morning” (B.Swinkels)
In case of Moving lines TEMPERATURE probes are the first to be looked at
because:
  Many noisy devices have temperature driven cycles (i.e. AirConditioners)
  Noise disturbances (EM) produced by electronic devices are typically very sensitive to temperature.
Temperature of DAQ
room

## Slide 9: … follow up: source identified
Changed on purpose DAQ room temperature by switching off air conditioner
Magnetic sniffing and Tapping on Electronic boards
Tapping on only this chassis at ~4 Hz generates ~4 Hz sidebands on many of the frequencies that responded to the previous temperature test
GPS synchronized 10 MHz timing generator

## Slide 10: … but “brute force” works better
NonNA (Non-linear Noise Analysis) by Francesco di Renzo, VIR-0406A-18
BUFFALO - Brute-force Utilities For Finding Annoying Lines and Others by Bas Swinkels, VIR-0420A-18
Example of Hrec line moving btw 650Hz and 700Hz. Correlated to Temperature of the electronics which demodulates photodiodes signals
“brute force” approach = search noise correlation in ALL slow channels (i.e. O(10000 channels)!)

## Slide 11: Do yourself…. correlations
Maddalena’s, 46852 “PRCL noise between 5Hz to 50Hz correlates with low frequency seismic noise”

## Slide 12: Do yourself…. correlations
Maddalena’s, 46852 “PRCL noise between 5Hz to 50Hz correlated with low frequency seismic noise”
SEISM band rms
PRCL asd amplitude
Several scripts exist, either in MATLAB or Python
…. ask commissioners 

## Slide 13: BruCo: Brute force Coherence
BruCo – computes and ranks coherence between Hrec and all AUXILIARY channels
(G.Vajente, https://dcc.ligo.org/LIGO-G1500230)
BruCo VIM plots: daily and ARCHIVE https://vim-online.virgo-gw.eu/?config=28

## Slide 14: BruCo plot
COHERENCE
Blue: LSC_DARM
Red: Projected coherence
LSC_DARM*COHERENCE
Dotted line:
Level of coherence
of pair of random signals

## Slide 15: … follow up: locate the noise source
Comparing the peak’s amplitude in all ENV accelerometers in CEB (26) helps locating the source
Example: 391Hz spectral line in all CEB accelerometers
… done by hand 
Nice to have utomated tool to do this

## Slide 16: Sniffing ….
Located the source of 391Hz line - Noisy cooling fan of turbo pump driver
Test accelerometer attached to chassis (double-side tape, always at hand )  coherent line
Used some paper to temporarily block the air outlet so to perturb the fan speed
Noisy fan because of ageing has been replaced and added vibration dampers
(VACUUM team - O3)

## Slide 17: Lines
Lines DB : archive of identified LINES https://apps.virgo-gw.eu/lines (by Gary Hemming). In charge: Neha Singh nsingh@astrouw.edu.pl
NoEMi V2.0 - LINES FINDER https://apps.virgo-gw.eu/noemi/?c=1. Ornella.Juliana.Piccinni@roma1.infn.it , gary.hemming@ego-gw.it

## Slide 18: Scattered light
A few slides tutorial:
What is it
How to recognize it … from spectral features (!)
Investigation techniques to locate the scatterer

## Slide 19
Credits: Antonino Chiummo
VIR-0013A-19

## Slide 20: Scattered light, some formulas
Frequency of the induced noise is proportional to the absolute value of scatterer velocity
Scatterer = some surface ground connected
in view of ITF beam (typical: optics on not suspended bench, vacuum chamber…)
Seismically excited: displacement x(t)
“ITF”
From: S.Braccini – VIR-0277A-09
Reference: Vinet J-Y, Brisson V and Braccini S 1996
Scattered light noise in gravitational wave interferometric
detectors: Coherent effects, PRD 54,2 1276
“G”
x(t)
x0 + x(t)
A change of x(t) of λ/2 means a change of optical path length of λ and a full rotation of the stray light vector
Frequency of noise = rate of rotation:
>>
NON LINEAR !
Expect non linear effects
(up-conversions) if
x : λ/2 = φ : 2π
φ = 4π x / λ
dφ/dt
2π
= 2 (dx/dt) / λ

## Slide 21: Scattered light – ARCHES
Example: shaking EIB with 0.5Hz sinus (46720)
.
NOTE that:
Height of arches is proportional to scatterer velocity
(i.e. if fmax = 20Hz  Vmax= 20*λ/2 = 10 λ = 10μm, a LOT!!)
Spacing of arches = half of scatterer period or 1/(2*fscatterer)
Second order of arches: scattered beam does double bounce
Blue line is
Bench displacement in microns:
Derivative of bench displacement

## Slide 22: Scattered light – SHOULDERS
Is just another way to look at the arches….
.
NOTE that:
Height of shoulder is proportional to “G”,
i.e. amount of scattered light which is coupling to ITF
Fitting the shoulders  measure G
Averaged FFT

## Slide 23: Measure and project
Use this G and x(t) in quiet condition
to evaluate the “expected contribution”
Of scattered light from that bench
Apply a known (some microm) motion x(t)
to the bench (use voice coil actuators)
Do fit of the shoulder to extract “G”
This technique is used to measure and monitor the
Amount of scattered light from benches
See Michal Was’s Noise Budget training session

## Slide 24: Scattered light - BUMPS
Case of light scattered from PR chamber blind flange:
Mono-chromatic vibration due to a fan or resonant mode of a mount or bench
(typically in the 10Hz to 100Hz range) is seen as a BUMP in Hrec (!)
Flange had vibration at 65Hz of ≈ 10^-9m
Small… BUT flange overall displacement was actually of O(microns) because of low frequency microseism
See Antonino Chiummo
VIR-0013A-19
Accelerometer attached to the flange

## Slide 25
One last feature:
when microseism increases bumps get shorter and fatter
all these features can be easily reproduced with a simple simulation, feeding one realistic x(t) in the formula
Try yourself:

## Slide 26: Scattered light hunting techniques
Inspect optical bench with infrared camera or IR cards to locate stray light beams
Tap the optics  DARM noise “explodes” when touching the critical one
MITIGATIONS:
add beam dumps/black glass to absorb scattered beams
Better align the beam on the bench, improve ITF global alignment:
SDB1 – DET team

## Slide 27: Elog 45006: ESQZB - Beam Dump Test and IR Viewer/Card Inspection (S.Chua et al)
Critical scattering
mitigated: improved alignment of SQZ injection, reduced OFI reflected power
3. Block the beam on bench following backwards the beam path and see when noise disappears
Scattered light hunting techniques

## Slide 28
White noise injected with shaker on SDB1 chamber wall
Cause is: light scattered off SDB1 onto chamber walls
Disadvantage: shaker injection is not localized, sometimes we prefer using instrumented hammer or “tapping”
SDB1 chamber
Scattered light hunting techniques
4. Shake or TAP (gently!) and see if some existing structures in Hrec grow….
(see Federico’s talk – ENV training Part 2)
Hrec
ACCELEROMETER

## Slide 29: Concluding …
Noise Hunter skills:
  Curiosity
  Perseverance
  Dedication
  …
  Curiosity!

## Slide 30: END
Thank you for your kind attention!